OPA Method for Measuring Submicromolar Ammonium

Updated: May 26, 2019 by BNW

References

ref: Holmes et al. 1999, Meeder et al. 2012 (incl. supplement)

Supplies:

Chemicals: sodium sulfite, sodium tetraborate, ethanol, opa

Lab Equipment: Fluorometer, 2 or 4 L dark bottle, repeat pippetter, quartz cuvette, something to rinse cuvette, kim wipes, gloves, sample tubes, 1 L grad cylinder, 150 ml grad cylinder, 10 ml grad cylinder, balance, stir bars, aged sample tubes.

Fluorometer: excitation 360 nm, emission 420 nm, low sensitivity.  Turner makes a fluorometer module specifically for this method, but you can use any fluorometer that allows you to adjust the wavelengths.  You may need to adjust the sensitivity differently on different instruments.

Reagents:

Makes ~ 2 L.

Always use fresh MQ and very clean glassware.

1. Sodium Sulfite Solution (0.8%)

Add 1 g sodium sulfite (sigma S-4672) to 125 ml MQ.

Stable for ~1 month at room temperature.

2. Borate Buffer Solution (4.0%)

Add 80g sodium tetraborate (Sigma S-9640) to 2 L MQ.

Stir to dissolve (takes awhile).

Store at room temperature.

3. Orthophthaldialdehyde (OPA) Solution (4.0%)

Add 4 g OPA (Sigma P-1378) to 100 ml high grade ethanol (Carlo Erba, Fisher).

Stir (swirl bottle) to dissolve (takes awhile).

Store in the dark at room temperature.

4. Working Reagent (WR)

In a large (>2 L) brown polyethylene bottle,

Mix 2 L of 1) 1.89L borate buffer solution, 2) 10 ml sodium sulfite solution, and 3) 100 ml OPA solution.

Allow WR to "age" 1 or more days before using.

Stable for 3 months when stored in the dark at room temperature.

Notes: For cruises, bring extra reagent and extra tubes because they have to age so you can’t improvise and whip reagent up on the fly.

Standards1

Primary Standard: 100 mM. Dissolve 0.52 g NH4Cl in 100 ml MQ.

Secondary Standard:  100 uM. Dilute the primary standard 1000x with MQ, i.e. 50ul primary in 50 ml.

1. Choose the range of standards appropriate for your samples.  For oceanic samples, 0.012 - 2 uM is appropriate.

2. Standards should be prepared in duplicate, and BF and 0 should be prepared in triplicate.

3.  Put 10 ml of the appropriate sample matrix into aged tubes2, and add the volumes given in the Table below to make standards.

The sample matrix must match the samples. Ideally, you would choose a sample with a low ammonium concentration prepare the standards using that sample, and then correct for the ammonium in that sample.

For seawater samples, use a sample with very low (<20 nM) ammonium to make standards. This can be either a deep sample (anything well below the nitracline, so 200 m or deeper) or a surface sample, if you know surface ammonium is <10 nM.

For rain samples, use MQ to make standards. NH4 may be 50-200 nM in MQ, so you may need to correct for this using the BF.

For estuarine samples, matrix effects may be variable between samples. I (BNW) have not run any estuarine samples using this method, so I would suggest testing the effect of the matrix by preparing multiple standard curves using different estuarine samples and comparing the slopes.  Alternatively, you can use the method of standard additions on each sample (Holmes, et al, 1999).

Table 1: own standard range

| [nh4] uM | Volume matrix (LNSW, MQ, etc) (ml) | Volume Secondary Standard (ul) |
| --- | --- | --- |
| 0.00 | 10.000 | 0 |
| 0.01 | 9.999 | 1 |
| 0.02 | 9.998 | 2 |
| 0.05 | 9.995 | 5 |
| 0.10 | 9.990 | 10 |
| 0.25 | 9.975 | 25 |
| 0.50 | 9.950 | 50 |
| 1.00 | 9.900 | 100 |
| 2.00 | 9.800 | 200 |
| 4.00 | 9.600 | 400 |
| 8.00 | 9.200 | 800 |

Samples and Laboratory Procedures

1. Decide if sample filtration is appropriate.3  Filter if necessary.

2. To collect sample, rinse aged tube2 with sample water twice and fill to 10 ml.

3. Store in the fridge until analysis.

4. Bring all samples, standards, and WR to room temperature.

5.  Turn off the lights during reagent spiking.4

6.  Dilute samples that you suspect will be > 8 uM.

7. Add 2.5 ml aged WR to each sample/standard.  Invert 5 times and store in the dark4.  Repeat for all.

8. Put BF standard in the dark with samples WITHOUT adding WR.  This will be your reagent blank.

9.  Incubate samples and standards at room temperature in the dark4 2-8 hours.  It is possible to read the samples up to 60 hours after spiking, but the fluorescence, and therefore the sensitivity of the method, decreases after 8 hours.

10.  Turn the fluorometer on a few minutes before running samples and check that the excitation and emission wavelengths are set correctly.

11.  Rinse the 1 cm cuvette with MQ.

12.  Add 2.5 ml WR to the BF sample, transfer to the cuvette, and read immediately.

13. Use a disposable pipette to transfer each standard and sample to the fluorometer.  It is not necessary to rinse the cuvette between samples unless you have just read a very high value.

14. Read one set of standards at the beginning of the run, read the samples, and then read the last set of standards.  This way you can correct for instrument drift.

15.  After you are done running samples, rinse the cuvette with MQ.  Turn off fluorometer.

Optional Filtration3 (not for oligotrophic seawater)

Materials: 0.2 um membrane filter (DO NOT use a filter made of glass or containing nitrogen. DO NOT use a cartridge filter. ONLY USE a membrane filter), 47 mm swinnex filter holder with adaptor, plastic tubing, peristaltic pump (optional), 2 L container Use all plastic in the setup- NO GLASS.

1. Assemble the swinnex with tubing, adaptor, and filter.  Use a zip tie to secure the tubing to the adaptor if using a pump.

2.  For a new filter, pump at least 2 L sample (or full strength ASW) through the filter before collecting samples.  This step is a pain, but DO NOT skimp on the rinse.  You will regret it when your samples are contaminated! I use a 2 L container to collect the filtrate so I know I rinsed at least 2 L.

3.  Pass the sample through the filter, rinse the tube, and directly fill the aged tube to 10 ml. Minimize sample contact with the outside of the tubing. If collecting from a Niskin bottle, you can attach the tubing right to the Niskin.

4.  Before collecting the next sample, pass enough water through the tubing and filter to rinse the entire thing twice.

5. Collect the next sample.

6. Repeat steps 3-5 until the filter clogs.  Then repeat step 2 and continue filtering.  Filters can be re-used a lot before they clog. They can be reused for multiple days. Don't rinse with MQ. Just keep rinsing with SW.

Calculations

General: Use BF to correct for contamination in the standards.  BF is the actual zero ammonium value.  Calculate concentrations from the BF-corrected standard curve. See Brittany for a spreadsheet.

Notes

1. For both standards and samples, this recipe uses 10 ml sample.  You then add reagent in a 1:4 reagent:sample ratio.  So, for 10 ml you add 2.5 ml reagent.  Other volumes can be used if needed, however smaller volumes should not be used for small concentrations. Smaller volumes magnify contamination.

2. All samples and standards should be prepared in "aged" tubes.  Aged tubes have been stored with either a small amount of WR in them or previously reacted samples for at least 1 week (the longer the better).  Over time the reagent removes all NH4 from the tube.  After using a set of tubes to run samples, simply leave the unused sample+WR in the tubes and save the tubes for next time.  Over time, some types of tubes (cheap Falcon tubes) can break down from the reagent; beware cracked caps.

3.  Filtration can contaminate ammonium samples substantially.  Ideally, samples should be run immediately after collection without filtration.  If biomass is really high, there are lots of particles, the samples must be refrigerated for more than 24 hours (for samples with low turnover rates only), or the samples must be frozen, filtration is unavoidable.  In these cases, follow the Optional Filtration steps above. Never filter oligotrophic seawater.

4. When spiking reading samples, keep the lights off in the room. it doesn't have to be pitch dark, but try to minimize light exposure. Keep the samples in a small box or cooler and take them out only to read.  While incubating, keep them very dark.  I have had problems when they're stored in something like a cardboard box that doesn't keep them completely dark.  I recommend putting them in a box inside something else (a cooler or drawer).

5. The background fluorescence (BF) sample is basically the blank for this method.  Because the BF has not had time to react with reagent, it reflects fluorescence due to the matrix and reagent and does not reflect the ammonium contamination possibly present in the standard matrix.  The BF should always be the same as or lower than the 0 standard.  If it isn't, the standards may not have been incubated in absolute darkness.